A. OBRIADINA, J.H. MENG, T. ULANOVA, A. BURKOV, H.A. FIELDS, Y.E KHUDYAKOV

Centers for Disease Control and Prevention, Atlanta, Ga, USA;

Diagnostic Systems, Nizhniy, Novgorod, Russia;

Southeast University School of Medicine, Nanjing, China

A NEW ENZYME IMMUNOASSAY FOR THE DETECTION OF ANTIBODIES AGAINST THE HEPATITIS E VIRUS

A new enzyme immunoassay (nEIA) for the detection of IgG or IgM anti-HEV activity using two new recombinant proteins has been developed. One of these proteins, designated as pB166, is located at 452-617 aa from within open reading frame (ORF) 2 and models the neutralizing epitope. The second protein, designated as MA-II, is an artificial second generation mosaic antigen composed of several short antigenic regions from within ORF2 and ORF3. The nEIA was evaluated using 3 panels: 1) acutely HEV infected patients (n=81), 2) non-A, non-B hepatitis (NANBH) (n=92), 3) normal blood donors from a non-endemic region (n=288). The nEIA detected IgG anti-HEV in 100% of acutely infected patients. When normal blood donors were tested more than 15% of the specimens were found IgG anti-HEV positive. To assist in determining the true value of these

specimens, they were retested using a large set of overlapping 30-mer synthetic peptides (n=71) spanning the entire ORF2. Ninety percent of these sera immunreacted with an average of 15 synthetic peptides. The remaining 10% did not immunoreact with any. These data suggest that the nEIA is at least 90% sensitive in detecting a remote HEV infection in normal blood donors residing in a non-epidemic region. The NANBH panel was first tested for IgG anti-HEV activity by 2 commercially available assays, EIA-1 and EIA-2. Concordance for IgG detection between the nEIA and the commercial kits was 73.3% and 75.3%, respectively. Concordance for IgM detection between the nEIA and a kit manufactured by GeneLabs was 97, 9%. In conclusion, a new EIA was developed for the detection of IgG or IgM anti-HEV activity in sera. A relatively poor concordance for the detection of IgG anti-HEV activity between nEIA and the commercially available assays requires further study.

American Society for Virology, 20th Annual Meeting, Madison.- USA, 2001.- P.169 (Р.5-2)