A. OBRIADINA, S. KLIMASHEVSKAYA, M. GLADYSHEVA, V. PUZYREV, A. BURKOV,

T. ULANOVA

Nizhniy Novgorod, RUS

A NEW ENZYME IMMUNOASSAY FOR THE DETECTION HERPES SIMPLEX VIRUS TYPE 2 – SPECIFIC ANTIBODIES

Objectives. The purpose of this study was to develop enzyme immunoassay (EIA) for the detection of IgG anti-HSV-2 activity using two new recombinant proteins as antigenic targets, and to evaluate these EIA with the aid of statistical methods.

Methods. Fragments of glycoprotein G (gG-2), comprising residues 525 to 578aa of herpes simplex virus type 2 (HSV-2) and glycoprotein D of HSV-2 (gD-2(266-394aa)), were expressed in the E. coli as GST fusion proteins to develop an assay for the detection and HSV-2 type-specific antibodies.

Results. A new enzyme immunoassay for the detection of IgG anti-HSV-2 (IgG-EIA) in sera was developed using two new recombinant proteins. The IgG-EIA was evaluated using serum specimens obtained from patients with culture-proven HSV-2 infection (CP) (n = 13) and from normal blood donors (BD) (n = 629). All specimens were additionally tested for IgG anti-HSV-2 activity by two commercially available EIAs. This new IgG-EIA detected anti-HSV2 activity in all specimens from HSV2 infected patients. When BD were tested the overall concordance between these three assays varied between 39 and 63.6%, concordance between positive samples ranged from 18.4 to 46.7%. In the absence of a gold standard the accuracy of these EIAs was assessed by the computer program based on a maximum likelihood approach using a ‘latent class’ model. This analysis estimated the IgG-EIA sensitivity and specificity to be within the range 98–100% and 95–100%, respectively.

Conclusion: The results show that the new two proteins-based enzyme immunoassays may be useful tools for the detection of type-specific HSV-2 antibodies. However, if only one assay is performed, careful interpretation of the results is indicated and for determination of the definitive HSV-2 serostatus, statistical models may be necessary.

14th European Congress of Clinical Microbiology and Infections Diseases.- Prague, 2004.-Р.844